| Name | Security Level |
| --- | --- |
| BELLSING® Hornet Series AMBA Receiver | Open |
| Version | A Total of 4 Pages |
| REV B |  |

BRC305C90100

Product Data Sheet

| Prepared by: | Sam Chen |  | Date: | 2020-09-22 |
| --- | --- | --- | --- | --- |
| Auditor: |  |  | Date: |  |
| Approver: |  |  | Date: |  |

Shenzhen Bellsing Acoustic Tech. Co., Ltd.

BRC305C90100 AMBA receiver

  - Description

Miniature balanced-armature receiver with stainless steel case for using in CIC (Complete In the Canal) and IIC (Invisible In the Canal) applications with standard response, can also be used in commercial hearing devices.

| Name: | Hornet Series AMBA Receiver |
| --- | --- |
| Part Number: | BRC305C90100 REV B |

  - Construction Specification

Weight:  0.09 g

Mechanical dimension：please refer to the drawing below

*A positive voltage applied to the negative terminal (­) will result in an increase in pressure at the sound outlet.

  - Features & Benefits

- Perfect package size for CIC, IIC applications

- Ideal choice of tweeter

- Outstanding efficiency

- Substantial maximum output

- Optimized mechanical shock resistance

  - Electro-acoustic Specifications

Test condition

- Acoustic tubing: 10.0 mm of 1.0 mm diameter tubing

- Acoustic coupler: IEC2CC coupler

- Constant voltage drive: 0.211 V RMS

- Environment: 23 ℃, RH 50%

Acoustic Parameter I

| Test Item |  | Nominal | Tolerance | Comments |
| --- | --- | --- | --- | --- |
| Sensitivity (dB) | @ 200Hz | 94.0 | +/- 3.0 |  |
|  | @ 500Hz | 94.0 | +/- 3.0 |  |
|  | @ 1000Hz | 94.0 | +/- 3.0 |  |
| Peak 1 | Frequency (Hz) | 3350 | +/- 350 |  |
|  | SPL (dB) | 103.0 | +/- 3.0 |  |
| Peak 2 | Frequency (Hz) | 6000 | +/- 500 |  |
|  | SPL (dB) | 98.5 | +/- 3.0 |  |
| Maximum output @ peak frequency (dB) |  | 125.0 | N/A | THD<10% |

Acoustic Parameter II

| Test Item |  | Nominal | Comments |
| --- | --- | --- | --- |
| THD (%) | @ 1/2 1st peak | < 5.0 |  |
|  | @ 1/3 1st peak | < 5.0 |  |

Electric Parameter

| Test Item |  | Nominal | Tolerance | Comments |
| --- | --- | --- | --- | --- |
| Impedance (ohm) | @ 500Hz | 106.0 | +/- 15% |  |
|  | @ 1000Hz | 118.0 | +/- 15% |  |
| DCR (ohm) |  | 100.0 | +/- 10% |  |

Nominal Frequency Response Curve

  - Package and Storage Condition

- Package: PET bubble pack, 1 pcs per bubble, 100 pcs per pack

- Storage temperature range: -40℃ ~ 60℃

- Storage humidity range: <60%